## ADS122xCalclulator
	ADS1220	http://www.ti.com/lit/ds/symlink/ads1220.pdf
	ADS122U04	https://www.ti.com/lit/ds/symlink/ads122u04.pdf
	ADS122C04	http://www.ti.com/lit/ds/symlink/ads122c04.pdf
	AppNotes	http://www.ti.com/lit/an/sbaa235a/sbaa235a.pdf
		http://www.ti.com/lit/an/sbaa201/sbaa201.pdf			Low-Pass Filter Design Considerations in Ratiometric Measurements (chapter 2)

Min/Max/Values	ADS122xx Calculations
2.3-5.5V	AVDD	4	V
-2.8<->0.3 DGND	AVSS	0	V	Negative AVSS increases measured temperature range.
0.01/0.05/0.1/0.25/0.5/1.0/1.5	Idac	1	mA	Recommended IDAC : pt100 – 1mA, pt500 – 0.5mA, pt1000 – 0.25mA
	Rref	2500	Ohms
	RTD 0°C resistance	100	Ohms
	Vref	2.5	V
	Min temp	-20	°C	Vin min	0.092160	Pt min Resistance	92.160	Min Resistance
	Max temp	140	°C	Vin max	0.153584	Pt max Resistance	153.584	Max Resistance
	Vcm = AVDD/2	2	V
1,2,4,8,16,32,64,128	Gain	16.3
	Selected Gain	16	Selected Gain	PASS
	G*Vin min	1.474565
	G*Vin max	2.457349
	PGA Common-Mode Voltage Requirements, 8.3.2.1
	When PGA is enabled				Limit for pass:
	Vcm min rule 1	1.428674	V	PASS	1.474565	G*Vin min
	Vcm min rule 2	1.000000	V	PASS	1.474565	G*Vin min
	Vcm max	2.571326	V	PASS	2.457349	G*Vin max
	Bypassing the PGA,, 8.3.2.1
	When PGA is disabled				Limit for pass:
	Vcm min	-0.100000	V	PASS	1.474565	G*Vin min
	Vcm max	4.100000	V	PASS	2.457349	G*Vin max
	Low Pass Filters : Corner Freq matching
	Input signal		R Reference
	R1 R2, Ohms	6200	R3 R6, Ohms	4990
	C1 C3 common, nF	4.7	C4 C6 Ccommon, nF	4.7
	C2 diff, nF 10*Ccomm	47	C5 diff, nF 10*Ccomm	47
		271.210	F, Hz	339.478	F, Hz
		269.883	F, Hz	334.333	F, Hz
	Self heating calculator
	Self-heating must not exceed 25% of the sensor’s tolerance specification!
	P reference	2.50	mW
	P rtd	0.10	mW
	Class A °C accuracy	0.2	°C
		0.05	25% Max °C
	Maximum self heating	0.50	°C/mW Max

## Accuracy
	Pt100
	Temp, °C	-20	0	30	100	150	200
	Class AA accuracy, °C Only for range 0-150 °C	N.A.	0.1	0.151	0.27	0.355	N.A.
	Class A accuracy, °C	0.11	0.15	0.21	0.35	0.45	0.55
	Class B accuracy, °C	0.2	0.3	0.45	0.8	1.05	1.3

	ADS122xx 1 Year Accuracy Estimation Calculator
	Ambient temp +/-	5	°C
	With Ref Resistor		Worst case with S102K				Best case with VH102K
	ADC INL		15.0	ppm	Max		6.0	ppm	Typical
	ADC Gain TC, PGA < 32, ppm/°C	2	10.0	ppm	Max	0.5	2.5	ppm	Typical
	ADC Offset TC , (uV/°C) /Vref → ppm/°C	0.24	1.2	ppm	Max	0.04	0.2	ppm	Typical
	Rref shelf life, 1 year		25.0	ppm	Max		5.0	ppm	Typical
	Rref TCR, ppm/°C	2.5	12.5	ppm	Max	1	5.0	ppm	Typical
	Total accuracy in ppm (SQRT)		33.3	ppm			9.6	ppm
	Total accuracy in Ohms for R=0°C		100.0033	Ohms			100.0010	Ohms
	Total accuracy in °C		0.0085	°C			0.0025	°C

	Without calibration in two fixed temperatures, the thermometer accuracy = pt100 Accuracy + ADS122xx Accuracy
	0°C ↔ 30°C		Accuracy				Accuracy
	pt100, class A, 30°C		0.21	°C			0.21	°C
	ADS122x		0.0085	°C			0.0025	°C
	Total		0.2185	°C			0.2125	°C
	Times better accuracy ADS122x vs. pt100, class A, 30°C		24.7				85.4

	In this case, S102K is good enough, because pt100 class A accuracy is much greater than ADS122x accuracy with S102K or VH102K
	Here the accuracy limiting factor is the pt100 nominal accuracy. The only way to improve this is to calbrate pt100+ADS122x with two fixed temperatures

	With calibration in two fixed temperatures (ice and gallium), the thermometer accuracy can be significantly improved in the range 0°C ↔ 29.76°C
	Uncertainty for 0°C can be less than 1m°C if method described in [1] is used
	Uncertainty for 29.76°C depends on Gallium impurity. It can be less than 13.7m°C (4N grade Gallium) if method described in [2] is used

	Gallium impurity	99.99	%	x2	0.0001	Taken from NIST Special Publication 260-157
				A	0.00732
	Estimated impurity uncertainty	0.0137	°C

	If Gallium with impurity 99.99% (4N) is used, ADS122x with S102K is still good enough (8.5 m°C vs. 13.7m°C)
	If Gallium with impurity 99.999% (5N) is used, ADS122x with VH102K will be better (2.5 m°C vs. 1.37m°C)
	Here accuracy limiting factor is the ADS122x accuracy and used reference resistor. Calibration of ADS122x with high purity Gallium (6+N) or triple water point does not make any sense.
	Note: uncertainty due hydrostatic-head, immersion (Heat Flux) and gas pressure are not taken into account due their low value in comparison with uncertainty of the 4 or 5 N grade Gallium and ice point. But if 7N grade Gallium and triple point water are used, they should be taken into account.

[1]	NIST Technical Note 1411	Ice, 0°C
[2]	NIST Special Publication 260-157	Gallium, 29.76 °C

## Calibration
Calibration only for ADS122x using external calibrated resistor

The example bellow will use a single resistor with nominal value of 100Ohms 0.1% tolerance,calibrated value 99.93295
Rnom = 100Ohms
Calibrated value of the resistor	99.93295	Ohms
ADC read value during calibration	5392449

1. Calculate multiplier	M = 1/((ADC/(223*Rcalibrated))
	155.457815889144

Calibration using two fixed point cells – Ice (0°C) and Gallium (29.76°C)
TBD

## Measurment
Using calculated value in Calibration tab, the measured temperature value can be found by the following formula:

	Example data
	M	155.457815889
	ADC	6000000
	A	0.0039083
	B	-5.775E-07
	Ro	100	Ohms

1. Calculate Resistance	R = (ADC/223)*M
R	111.192	Ohms

2. Calculate Temperature	T= ( -A+SQRT(A2-4*B*(1-(R/Ro))) ) / (2*B)
T	28.759	°C